## Slide 1: An Improved Correlation Instrument for Measurements of Noise in Amplifiers and Transistors
Andrew Fox
Mentor: Dazhen Gu

## Slide 2: What are we measuring?
Noise Parameters in the 1-2 GHz regime
Noise Figure (NF): the amount of noise that an active device adds to a signal (in dB)
  Minimum Noise Figure (Fmin): lowest achievable NF for a device
  Reflection coefficient (Γopt): optimal reflectance for lowest noise (complex)
  Noise Resistance (Rn): how sensitive the NF is to deviations from the optimal reflectance
NF is equivalent to Noise Temperature (NT)
Accurate measurements are vital for low-noise amplifiers
More difficult to maintain precision as NF gets smaller
Notes: talk about motivation of communication devices on this slide

## Slide 3: Source-Pull Method
Send a known amount of noise with varying source impedance into amplifier, analyzer measures noise figure/noise power
Tuners are mechanical, and often large
At best fraction of a decibel precision
BSW Testsystems & Consulting AG. Impedance Adaptations / Noise Parameter.
Notes: Talk about deficits of mechanical: can break, have to sweep all of the impedances (slow). also have to be so large
about half a decibel

## Slide 4: Correlator Method
Directly samples noise waves emanating from input (Ch 1) and output (Ch 2) ports of the DUT
Uses an Analog to Digital converter (ADC) and digital signal processing to convert time series data to frequency series and find power density
Processed data used to calculate noise parameters (and NF)

## Slide 5: Correlator Method - Components
Notes: adapted from a previous publication, get away from variable noise due to precision

## Slide 6: Redesigned Structures
Greatly reduced size ( < 50% original size)
Added adjustable attenuator, second stage amplifier
CH 1+2
Smaller structure
.5” movement in two axes
DUT
Notes: Say I

## Slide 7: CAD Design
Finalized Design
Ensures movement in three axes for DUT of different sizes
2.5”
.5”
2.5”
.5”

## Slide 8: Assembled Device
Parts ordered, delivered, and assembled without incident or delay
S-parameters measured to help calibrate device

## Slide 9: Measurement Steps
S-parameters – find the gain of both channels
Continuous Wave (CW) – precisely find gain vs frequency
Noise Diode (ND) – send a noisy signal with known NT through
DUT – Measure the DUT to extract noise parameters

## Slide 10: CW Tone Processing
Send known tones every 5 MHz from 1-2 GHz
With processing steps, can find the gain and phase difference
Repeatability across different powers

## Slide 11: ND Data Processing
Pilot signal – we know the noise emitted from it, and can find the gain of that noise
Still used for calibration
Very similar gain as found via CW

## Slide 12: Calibration Comparison
Very similar results (RMS error of .21 and .23 dB)
Some minor corrections still needed

## Slide 13: Conclusion
Correlator vs Source-Pull:
Reasonable preliminary output
Hoping for better
Design & Assembly:
50% original size
Improved dynamic range
Calibration:
CW tones + noise diode
Reproducible results
Future Steps:
Measure DUT
Known Gain -> measure NF

## Slide 14: Thank You
Special thanks to:
Dazhen Gu
Xifeng Lu
Dan Kuester
Paritosh Manurkar
CTL Div. 675

## Slide 16: Noise Figure & Noise Temp

## Slide 17: CW Calibration
CW Tones sent from 1-2 GHz every 5 MHz

## Slide 18: Noise Signal Calibration

## Slide 19: Phase Comparison
Phases of the two are virtually the same

## Slide 20: Components
Channel 2 (Out)
Channel 1 (In)
Both channels are read by an ADC

## Slide 21: Device Under Test (DUT)
Arbitrary low-noise transistor or amplifier

## Slide 22: Circulator
CF1020
17 dB isolation

## Slide 23: Isolator
CI1020
17 dB Isolation

## Slide 24: Low Noise Amplifier
FMAM 1019
40dB gain
.8 dB NF
11dBm output power

## Slide 25: Attenuator
PE70A1049S
0-41dB Attenuation
Allows for improved dynamic range

## Slide 26: Band Pass Filter
CFB 10201
Only allows 1-2 Ghz, 50 dB attenuation otherwise

## Slide 27: Medium Power Amplifier
ZX60-2534MA+
~ 40 dB gain
2.6 DB NF
18 dBm output power

## Slide 28: Low Pass Filter
VLFG-2000+
~42 dB attenuation above 2GHz